ATTACHMENT 1 REV 2
FULL CALIBRATION OF THE QRAE-3 using gas bottle

- Press and hold the [Mode] key to turn unit on.

- Allow unit to warm up. Select fresh air calibration and verify unit passes calibration.

- Ensure gas bottle is not expired.

- Hold down the [Mode] key and the [Y/+] key at the same time until the screen changes to [Enter Password].

- The default password is “0000”. Use the [Mode] key and the [Y/+] key to enter the password.

- Press the [Mode] key until OK is highlighted. Press the [Y/+] key to select.

- [Calibration] will be highlighted. Press the [Y/+] key to select.

- Using the [Mode] key, scroll down until Multi Span is highlighted. Press the [Y/+] key to select. The unit will ask you to connect the gas bottle.

- Turn on gas bottle and connect to QRAE-3. Unit will automatically begin calibrating, counting down from 60 seconds.

- After 60 seconds, unit will alarm, display readings and indicate pass/fail.

- If the unit indicates that a sensor has failed, repeat the calibration process. If a sensor fails a second time, remove the unit from use and replace the sensor.

- Using the [Mode] key, scroll across until Exit is highlighted. Press the [Y/+] key to select.

- Again, using the [Mode] key, scroll across until Exit is highlighted. Press the [Y/+] key to select.

- Unit will return to normal run mode and is ready to sample.

- Press and hold the [Mode] key to turn unit off.