ATTACHMENT 4 Rev 0
CALIBRATION OF THE REED INSTRUMENTS MODEL

TM-8811 ULTRASONIC THICKNESS TESTER

Operating Instructions

- Insert ultrasonic sensor into the probe connectors on the meter.

- Press the power button to turn the meter on

- The LCD will display “0” and the selected unit of measure (mm or inch)

Initial Calibration

- Apply a small drop of coupling gel on the 5mm calibration test block on the meter.

- Press CAL button and “CAL” will appear on the display.

- Place the sensor on the calibration test block while ensuring the coupling indicator ((•)) appears on the display and the red coupling LED turns on.

- Depending on the selected unit of measurement either “5.0 mm” or “0.197 in” and “CAL” will flash interchangeably on the display.

- While continuing to hold the sensor on the calibration test block, press CAL again to save.

- Calibration is now complete, and the ultrasonic sensor can now be removed from the calibration test block.

NOTE: Results from the calibration procedure will automatically save in the meter. It is not necessary to calibrate the meter each time the unit is powered on unless measurements appear to be inaccurate. The calibration procedure confirms both the meter and ultrasonic sensor are functioning properly.

Material Mode Selection

Selecting Pre-loaded Materials

The meter has been pre-loaded with 11 common materials to allow for quick setup and testing. In order to select the appropriate setting, follow steps 1 through 4 below.

- While the meter is on, press the SELECT button

- The display will either indicate “CdXX” or “XXX”. “Cd” stands for code and “XX” is the pre-loaded material number which ranges from 1 to 11 as shown below.

- “XXXX” is a 4-digit number that represents the sound velocity of the material that can be adjusted by the user (For entering and adjusting Sound Velocity please see “Setting the Sound Velocity for a known Material”).

Cd01- Steel		Cd05- Brass		Cd09- PVC

Cd02- Cast Iron	Cd06- Zinc		Cd10- Gray Cast Iron

Cd03- Aluminum	Cd07- Quartz/Glass	Cd11- Nodular Cast Iron

Cd04- Copper		Cd08- Polyethylene	Cd12- User Defined Sound Velocity

- Press the Up/Down arrows to select the desired material code. Press SELECT button to confirm your selection. The display will show “0” confirming the material has been selected. If you select a material code without confirming the selection, the meter will automatically resume normal operation after a few seconds.

Measuring Procedure

Once the material has been selected, follow steps 1 through 4 below to perform the measurement.

- While the meter is on, press the mm/inch button to select the correct unit of measure.

- After applying coupling gel on the material, place the sensor firmly against the desired measurement area. NOTE: For most applications, a single droplet of coupling gel is sufficient

- Verify that the Coupling indicator ((•)) appears on the display and the coupling LED light turns on

- The LCD will display the measured value and will hold this value until a new measurement is taken or until the unit is turned off.

NOTE: if the coupling indicator ((•)) does not appear on the display, or the measure values appear to be erratic, verify that there is an adequate amount of coupling gel in between the ultrasonic sensor and the material under test. It is also important that the ultrasonic sensor sits flat against the material.

Setting the Sound Velocity for a Known Material

If the material under test is not pre-loaded, the sound velocity of the material under test can be entered. To enter the appropriate sound velocity, follow steps 1 through 9 below.

- While meter is on, press SELECT

- The display will either indicate “CdXX” or “XXXX”

- Press the UP/DOWN arrows to scroll through the material codes until you see the 4-digit value

- This 4-digit number represents the last set sound velocity of a material

- Press the SELECT button to confirm your selection. The display will show “0” confirming the material has been selected. If you select a material code without confirming the selection, the meter will automatically resume normal operation after a few seconds.

- Press the VEL button and the display will show the last saved sound velocity measurement.

- The Sound velocity can be adjusted by pressing the UP/DOWN arrows to match the sound velocity for the material under test as shown below.

Aluminum- 0.250 in		Zinc- 0.164 in		Tin- 0.117 in

Silver- 0.142 in		Gold- 0.128 in		Steel, Common- 0.233 in

Steel, Stainless- 0.226 in	Brass- 0.173 in	Copper- 0.186 in

Iron- 0.233 in			Case Iron- 0.173 in	Lead- 0.094 in

PVC- 0.094 in			Titanium- 0.236 in	SUS- 0.240 in

Epoxy Resin- 0.100 in		Grey Cast- 0.180 in

*Other materials listed in owner’s manual*

- Press the UP/DOWN arrows to adjust the sound velocity value.

- Press the VEL button to confirm your selection. The display will show “0” confirming the sound velocity has been set. If you set a sound velocity without confirming the selection, the meter will automatically resume normal operation after a few seconds.